## Slide 1: Team 60: RADIAN Bi-Weekly Update 1 Matthew Blackwell, Mitchell Hoffman, Wyatt Johnson Sponsor: Dr. Arya Menon TA: Manoj Kumar Balamurugan
Notes: Michelle

## Slide 2: Problem Statement
Falls are the leading cause of injury in elderly
Fast detection saves lives through quick response
Current Solutions - Limitations:
Wearables → must be worn and charged
Push-buttons → user must be conscious/able to press
Camera → effective but invade privacy

## Slide 3: RADIAN System Summary
Proposed Solution:
Automatic Fall Detection
  Works in real time
  No user action needed
Privacy-Preserving
  No camera
  Protects independence
Reliable Response
  Immediate accurate alerts to caregivers

## Slide 4: System Overview
Battery
TI AWR6843
Raspberry Pi 3
PCB
AI Algorithm / Signal Processing
Phone App
Serial (UART or USB)
External Alarm
Watchdog Timer
Matthew Blackwell
Wyatt Johnson
Mitchell Hoffman
External Casing and Radome

## Slide 5: Major Changes for RADIAN
Data Subsystem | Physical Subsystem | Wireless Subsystem
Changing from a decision tree model to a Convolutional Neural Network (CNN), causing a redesign of Data Subsystem. | Redesigned the Power PCB | No major changes since 403
Notes: Eric

## Slide 6: RADIAN Timeline
Planning
Designing
Subsystem Testing
Integration
System Testing
System Demo
Data
Wireless
Physical
Switching ML models has caused the Data Subsystem to be shifted back into the design phase
Ongoing PCB testing
Improving code for radar to get more accurate data for integration
Notes: Eric

## Slide 7: Data Subsystem
Changes since 403 ~35 hrs of effort | Ongoing progress and plans until the next presentation
The previous ML model was not sufficient for 404, in both technical merit and reliability when it comes time to integrate Researched new ML models for RADIAN, landed on a CNN for reliability and ability to learn different features Documented how data frames will work | Code a baseline model that can be improved as we start to integrate Integrate with Wireless to communicate through the Gotify server Start baseline trials and data cleaning to train ML model Document process for recording trials

## Slide 8: Data Subsystem
Gather Data
- Scan a variety of falls (sitting, standing, laying down, etc.)
- Scan a variety of non-falls (standing still, sitting still, walking, sleeping, etc.)
Label Data
- Labeling each of the falls and non-falls and how they occurred
Clean Data
- Removing the outliers, unnecessary data, missing, and irregular data
Feature Selection
- What constitutes a fall?
- Acceleration, Velocity,
Jerk, Duration of Movement, Height Change, Posture after Event
Train Data
- Create a training set, validation set, and testing set of our data
Report Results
- Using a CNN, report results within a ±5% accuracy

## Slide 9: Data Gantt Chart Data’s Two Week Plan

## Slide 10: Wireless Subsystem Owner
Accomplishments since 403 ~20 hrs of effort | Ongoing progress/problems and plans until the next presentation
The code for the radar’s data processing can track an object (person) and provide x, y, z, and velocity data for it at each time frame through UART The mobile app has a basic UI for the caregiver with simulated test alerts | Improve radar’s data processing code to provide more accurate data Debug what the script outputs to have a more seamless integration with the ML model

## Slide 11: Wireless Subsystem
Radar is able to receive a configuration file and the code for the data processing produces positional data and velocity data with timestamps
The data processing needs to be more accurate and and output needs to be altered for easier integration with the ML model.
Mobile app can test simulated alerts, but the fall log needs debugging so it can log the alerts and accurately
Add more features for the caregiver

## Slide 12: Wireless Gantt Chart Wireless’s Two Week Plan
WIRELESS GANTT CHART | 01/12/26 | 01/19/26 | 01/26/26 | 02/02/26 | 02/06/26 | 02/16/26
Review the code for mobile app and radar to see where it was left off and decide what works still needs to be done and/or fixed. |  |  |  |  |  | 
Research ways to improve what is wanted/needed from radar |  |  |  |  |  | 
Debug and improve radar's data processing code |  |  |  |  |  | 
Begin to integrate with ML model and train with the radar's data |  |  |  |  |  | 
Test alerts going to the mobile app, debug/improve from there |  |  |  |  |  | 

## Slide 13: Physical Subsystem
Accomplishments since 403 ~40 hrs of effort | Ongoing progress/problems and plans until the next presentation
Assembled Rev 2 Power Battery PCB. Fully tested the buck converter using a bench power supply. Verified the battery protection circuit is not tripped under no-load conditions. Confirmed the buck converter operates correctly at no load while simulating two series 18650 cells. | Continue MP2615 charger bring-up. Validate full range of 2S2P operation (8.4–6.0 V) with stable 5V. Confirm S-8252 UV/OC trip + recovery. Write a brief test report (measurements, pass/fail, photos, issues).

## Slide 14: Physical Subsystem
Barrel Jack / 12V DC Input
MP2615 Charger IC
2S2P Li-Ion 18650 batteries
Battery Protection (S-8252)
5V Buck Converter (AP64501)
System 5V Rail (USB-C x2)

## Slide 15: Buck Converter Load Test (Vout vs Iout)
Current (A) | Voltage (V)
0.10 | 4.98
0.25 | 4.98
0.50 | 4.98
1.00 | 4.95
1.30 | 4.93
1.50 | 4.90
1.70 | 4.87
1.80 | 4.84
1.90 | 4.81
2.00 | 4.80
2.10 | 4.78
2.20 | 4.77
2.50 | 4.70
4.63

## Slide 16: Physical Gantt Chart Physical’s Two Week Plan

## Slide 17: RADIAN Validation Plan
Notes: Michelle

## Slide 18: RADIAN Gantt Chart RADIAN’s Two Week Plan
Notes: Michelle

## Slide 19: Thanks and Gig ‘Em Any Questions?